TRAINING PLAN AND REPORT OF AnaLight® 4D

By Juan Wang

NEURASYNC Academic-Industrial Training Network; Farfield Group Limited

Introduction

The AnaLight® 4D is an optical interface-analytical instrument for the determination of the thickness, density and mass of thin films captured on the surface of an AnaChip. It is important that the users familiarise themselves with safety and technical information for the correct operation and maintenance of the AnaLight® 4D if the equipment is to be operated correctly and to its best. A full training of the operation and software (AnaLight® DAQ, AnaLight® Explorer) for the start users should be given before the users can run the experiments by themselves.

The initial training covers operation and maintenance of the AnaLight® 4D, the on-line data acquisition software AnaLight® DAQ and the off-line data analysis software AnaLight® Explorer. The users may receive further training for AnaLight® Script Editor and Affinity/Association/Dissociation templates. AnaLight® Script Editor is a software package to allow either semi-automated or automated operation of the fluidic components within the Farfield Dual Polarisation Interferometer AnaLight® product range. Affinity, Association and Dissociation Templates are used for further affinity and kinetics analysis of experiment data.

Trainees

New customer/users of DPI, application specialists or distributors.

Training Plan

- Introduction of Principle of Dual Polarization Interferometry (DPI)

Time duration: 30 min presentation.

Targets:

Basic understanding of DPI principle and all parameters that can be obtained with DPI measurements.

- Instruction on AnaLight® 4D and Basic DPI Measurement Setup

Time duration: Half-day demonstration, hands-on training and practice.

Targets:

Get to know main components of the instrument (pump, valve controller, valves, fluidic system, manifold, and chip carrier);

How to power up the system, assemble and insert the chip carrier;

Set parameters of pump;

Make sure the fluidic system no leakage and blockage;

Other preparations for experiment.

Remark: will need some high quality water (Millipore or HPLC water); vacuum pump and sonicator for degasing water; gaskets and unmodified AnaChips.

- Training on AnaLight® DAQ

Time duration: Half-day demonstration and hands-on training; half-day practice.

Targets:

Start up a measurement;

Fringe selection; data acquisition settings and parameters;

Set up options (a filename, temperature, data acquisition rate, chip type, calibration materials);

Calibration steps;

Do injections and monitor /adjust instrument setting during an experiment.

Remark: 1 litre of high quality water degassed before use; 20 mL 80% ethanol/water by weight (analysis grade ethanol); unmodified AnaChips and gaskets.

- Training on AnaLight® Explorer

Time duration: Half-day demonstration and half-day practice.

Targets:

Open/convert and save data files;

Data display and navigation (customise the screen layout and traces to display; scale and scroll plots);

Calibration steps (linearization, chip calibration and bulk RI calibration);

Calculation of layer thickness, RI, density and mass;

Analysis of anisotropic layers;

Maths functions;

Data export and presentations (analysis tables; overlaid defined regions)

- Practice Experiments and Data Analysis by Trainees.

Time duration: one-day practice and questions

Targets:

The trainees should be able to run a simple demo experiment and do the data analysis. Help trainees with questions arising during the experiment.

Remark: Different demo experiments could be prepared according to the field of trainees. Buffers, water and 80% ethanol/water will be needed.

- Training on AnaLight® Script Editor and Affinity/Association/Dissociation templates. (optional)

Time duration: one-day demonstration and questions

Targets:

Get familiar with commands and edit experimental scripts

Run scripts automatically or manually

Get affinities and kinetics from binding curves

Conclusions

The training takes 3.5 to 5 days according to how much training covered and how many trainees involved. The first to forth step of training is usually necessary and most important However, if the trainees can get themselves familiar with basic concepts of DPI and its operation principle by reading support materials/documents suggested by the trainer, the training can be more effective and quicker. Hands-on training works better than presentations, therefore the presentation is only used for introducing principle of DPI. To give enough time/chances for each trainee to take part in the hands-on training and practice, a relative small group of trainees (below 5) works better than large groups. A list of frequently asked questions arising from the training (sent to trainees after the training) usually helps trainees to get better from the training.